Zaini Hamzah et al:   DISCRIMINATOR SETTING AND COCKTAIL PREPARATION FOR ANALYSIS OF ALPHA AND BETA EMITTERS IN AQUEOUS SOLUTION USING LIQUID SCINTILLATION COUNTER

The Malaysian Journal of Analytical Sciences, Vol 15 No 1 (2011): 27 - 36

[image: image1.png]



DISCRIMINATOR SETTING AND COCKTAIL PREPARATION FOR ANALYSIS OF ALPHA AND BETA EMITTERS IN AQUEOUS SOLUTION USING LIQUID SCINTILLATION COUNTER

(Penetapan Diskriminator dan Penyediaan Koktil Untuk Analisis Pemancar Alfa dan Beta Didalam Larutan Berakueus Menggunakan Pembilang Sintilasi Cecair)

Zaini Hamzah1*, Masitah Alias1and Zaharudin Ahmad2
1Faculty of Applied Sciences, 
Universiti Teknologi MARA, 40450 Shah Alam, Selangor, Malaysia 

2Radiochemistry and the Environment Group, 
Malaysian Nuclear Agency, 43000 Bangi, Selangor, Malaysia
*Corresponding author: zainihamzah@live.com
Abstract

Liquid scintillation counting (LSC) is not only being used to measure pure beta emitters, but it can be used to measure both alpha and beta emitters simultaneously. Measurement of alpha and beta emitters in aqueous solution is done using a single sample. For the sample preparation, colorless detergent or emulsifier was used to incorporate the water into an organic based scintillator to produce a clear homogeneous solution, since this is the best form to give the highest count rate and detection efficiency. The instrument also need some attention, where after calibration, the LSC was set for the discriminator level which is suitable for measurement of both alpha and beta radiations. In this study, the focus is on the development of the best scintillation cocktail and establishes the best discriminator setting. From this study the best proportion of scintillation cocktail is 2:4:4 for water, toluene, and Triton-N101 (emulsifier) respectively and the best discriminator setting for alpha and beta counting are 120. 

Abstract in English

Keywords: alpha emitter, beta emitter, cocktail, LSC, scintillant.
Abstrak

Pembilang sintilasi cecair (PSC) bukan hanya digunakan untuk mengukur pemancar beta tulen, tetapi boleh juga digunakan untuk mengukur kedua-dua pemancar alfa and beta serentak. Pengukuran pemancar alfa dan beta di dalam larutan berair boleh digunakan menggunakan sampel yang sama. Untuk penyediaan sampel, emulsifier tanpa warna telah digunakan bagi mengikat air ke dalam sintilator berasaskan organik untuk menghasilkan larutan homogeneous, kerana ini adalah merupakan bentuk terbaik yang memberikan kadar bilangan dan efisiensi pengesanan yang tertinggi. Peralatan juga memerlukan perhatian, dimana selepas proses kalibrasi, PSC perlu penentuan tahap diskrimanator yang sesuai untuk pengukuran kedua-dua radiasi alfa dan beta. Ini adalah untuk memastikan kadar limpahan dari kedua-dua radiasi rendah. Di dalam kajian ini, fokus ditumpukan kepada membangunkan koktil sintilasi dan menentukan tahap diskriminator yang sesuai. Dari kajian ini didapati kadar terbaik bagi koktil sintilasi adalah 2:4:4 bagi air, toluen dan Triton N-101 (emulsifier) masing-masing dan tahap diskriminator terbaik adalah pada 120 untuk pengukuran zarah alfa dan beta serentak.
Kata kunci: pemancar alfa, pemancar beta, koktil, PSC, sintilan.

Introduction

Measurement of alpha and beta emitters in environment especially water samples are normally carried out using separate method, since both radiations having different properties. Recently, with the advancement of liquid scintillation counting (LSC), the measurement of radionuclides in the natural environment where radionuclide concentrations are low, becoming more feasible since the instrument background is very low. The contribution of instrument background to the precision of the measurement is often important. The “low level” instruments have their background reduction features that are the efficiency to background (E2/B) factor is increased or minimum detectable activity (MDA) is decreased [1].

This method is based on the recently available liquid scintillation counting (LSC) systems with the ability of discrimination of alpha and beta emitters by the time distribution of the light emission each generates from scintillators, in the term of pulse shape discrimination (PSD) or pulse decay analysis (PDA). Pulse Shape Analysis (PSA) is a pulse shape discrimination technique that is based on a method that integrates the charge of ‘tail’ of scintillation pulse and compares it with the total charge in the same pulse. Different settings of the PSA level assign the pulse into either a long (alpha-like) or short (beta-like) category. Thus, different PSA settings allow pulses to be categorized according to their length (shape). Typically, increasing the PSA setting will direct more pulses toward the long or alpha category. However, depending on how the technique is implemented, the reverse can also be true, i.e., increasing the setting may direct more alpha counts into the beta category.  There are several methods of accomplishing pulse shape analysis including slow crossover timing, fast crossover timing, and constant fraction of pulse-height trigger [1].
Water quality is an important aspect of environmental studies. Natural waters contain both alpha (such as 238U) and beta (such as 40K) emitters in widely varying concentrations which are responsible for a generally small fraction of the total dose received from natural and artificial radioactivity [2]. For practical purposes, the recommended guideline for activity concentrations are 0.1 BqL-1 for gross alpha and 1 BqL-1 for gross beta in the drinking water [3]. The recommendations do not differentiate between natural and man-made radionuclides. Below these reference levels of gross activity, drinking water is acceptable for human consumption and any action to reduce radioactivity is not necessary. In the last decades, nationwide studies have been performed in practically every developed country, where drinking water is regularly sampled and analyzed. In developing nations, data on the natural radioactivity in water are not always available. According to a UNSCEAR report [4], drinking water is considered to be an important factor in increasing the natural radiation exposure in humans. 
The occurrence of natural radionuclides in drinking water poses a problem of health hazard, when these radionuclides are taken into the body by ingestion. Several naturally occurring alpha or beta emitting radionuclides such as 238U, 226Ra, 222Rn, 210Pb, 228Ra and others are frequently dissolved in domestic water supplies and their concentrations vary over an extremely wide range, mainly depending upon the amount of radio elements present in bedrock and soil with which the water comes in contact [5].
Terlikowska et. al. [6], has done some work on the application of alpha/beta discrimination in liquid scintillation counting for the purity control of 99mTc medical solutions. Shing-Fa Fang et al. [7], has worked on the comparison of Alpha/Beta separation performance of Commercially Available Scintillation Cocktails and counting by QuantulusTM 1220 Liquid Scintillation Counter. Maurizio Forte et. al. [8], has done some natural radionuclides measurements in drinking water using Ultra-low Level Liquid Scintillation Counting. Cantaloub et.al. [9] studied the interaction of sample, cocktail and headspace volume when measuring aqueous radon in small volume sample. The study focused on measurement of radon using commercial cocktail, Ultima Gold F. 

In this study, the focus is on the development of the best scintillation cocktail and establishes the best discriminator setting for LSC TRICAB 2700 from Packard that available in Malaysian Nuclear Agency (NM) in Bangi, Malaysia. The scintillation cocktail was developed using emulsifier (Triton N-101) mixed with solvent (toluene) and scintillator (2,5-diphenyloxazole, PPO and 1,4-bis(5-diphenyloxazol-2-yl)-benzene, POPOP). 
Experimental
Instrument Calibration 

Liquid Scintillation Counter TRICAB 2700 is a Packard instrument which is capable for measuring both alpha and beta radiations in water. Before any samples are counted, the instrument must be calibrated. This is performed automatically with the use of the calibration (SNC) protocol plug and the unquenched 14C calibration standard. 

Alpha/ beta Calibration and Discriminator Setting

Using Packard TRICAB models with alpha/beta discrimination, the optimum setting is the setting where there is equal and minimum spillover of alpha pulses into the beta Multi Channel Analyzer (MCA) and beta pulses into the alpha Multi Channel Analyzer (MCA). The determination of an optimum PDD requires two standards: one of the pure alpha emitter of interest and one of the pure beta emitter of interest. 241Am and 36Cl samples in Packard Ultima Gold AB, a cocktail specifically designed for alpha/beta separation were used for setting the percent spillover or percent misclassification of alpha and beta emitters. To achieve this, one need to vary the discriminator setting on the instrument (LSC). The instrument determines it optimum setting that results in the minimum misclassification of alpha and beta activity, and will generate the percent misclassification plot on demand. Several observations can be made by comparing these two curves. The first observation is the obvious shift in the instrument determined optimum PDD to a slightly higher value with TR-PDA than without TR-PDA. The next observation is the flattening of the alpha curve. 

Cocktail development

Cocktail development is the method where we attempt to incorporate the water in an organic solvent.  A triangular method was used to get the best proportion of water, organic solvent and Triton N-100 with 4.0 g/L PPO and 0.4 g/L POPOP (Figure1).
[image: image2.png]



Figure 1: The position of vials for different composition of water, toluene and Triton N-100
Results and Discussion
Instrument Calibration

During the calibration procedure, the high voltage applied to photomultiplier tube (PMT) is adjusted individually until the tubes are normalized (synchronized) in their response to the 14C calibration standard. Then, the high voltage to both tubes is adjusted simultaneously until the end point of 14C spectrum fall in the appropriate position on the spectra analyzer (on-board 400-channel) multi-channel analyzer.

It is highly recommended that the SNS cassette (containing the 14C calibration standard) be left in the instrument at all time. If the instrument is not IDLE, the calibration procedure is performed once every a 23-hour timer is checked. If 23 hour has elapsed since the previous calibration, the instrument will automatically perform the SNC procedure. If the time has NOT elapsed, the SNC cassette is bypassed. When the calibration is complete, the message SYSTEM NORMALIZED appears on the printout. This indicates that the instrument is ready for counting.   

Discriminator Setting

Separate optimum PDD setting should also derived for unique alpha and beta standard pairs. For gross alpha and gross beta measurements where the particular radionuclides may not be known, an alpha and beta standard of similar energy to the alpha and beta in the samples is desirable. To arrive at the optimum setting, each standard is counted individually at range of PDD settings and the percent misclassification of alphas into the beta MCA and vice versa are plotted against PDD on the same graph. When only the beta emitter is of interest, a PDD value below the instrument determined optimum may be used, which minimizes the misclassification of alpha activity into the beta MCA at the expense of reducing the beta efficiency. Similarly, when only the alpha emitter is of interest, a PDD value greater than the optimum, can be used. This minimizes the mis-classification of beta events into the alpha MCA at the expense of reduced alpha counting efficiency. 

The effect of Time-Resolved Pulse Decay Analysis (TR-PDA) on the misclassification curve is slightly higher value with TR-PDA and the flattening of alpha curve is the cause of the shift in the intersection of the curves and results in a higher optimum PDD value. The flattened shape of the alpha curve is due to TR-LSC discrimination of misclassified alpha events in the beta MCA. At this higher value, beta misclassification is also reduced. Manual adjustment the discriminator to a higher or lower value than the optimum will further reduce either beta or alpha misclassification at the expense of some loss of alpha or beta counting efficiency, respectively.

Figures 2(a) to 2(e) show the alpha/beta spillover curve for 36Cl/241Am. From these curves we have chosen the discriminator 120 with the percent spillover alpha:beta (1.45:1.63).

Cocktail development

In this study, the emulsifier used was Triton N-101 and the solvent was toluene scintillation grade. The amount of primary and secondary scintillator was 4.0 g/L and 0.4 g/L of PPO and POPOP respectively. Using the triangular method [10], 36 vials were used to be filled with different proportions of water, toluene and emulsifier (Triton N-100).

Tables 1(a) and 1(b) show the proportion of water, toluene, and triton N-101 in different vials. The appearance of liquid in each vial was recorded as cloudy, clear and opaque. The aim of this cocktail is to get a clear solution and that solution is giving the best count rate. The count rate of 3H (pure beta emitter) and 226Ra (alpha emitter) are listed in the column 2 of Table 1. 

Both 3H and 226Ra show the same characteristics when mixed with cocktail. The aim of this cocktail preparation is to get the clear stable solution, with the best count rate as well as giving the best merit value. From Tables 1 (a) and (b), the best proportion for the mixture of water:toluene:Triton N-101 is 2:4:4. So, for 20 mL mixture, the amount of water is 4 mL. 

[image: image3.png]Alpha/Beta Spilover Curve for:Clam.
Discinivateri11512% 5o Over ptarsets (147151

——A
—=—B /

Oerminator Setting





[image: image4.png]40

Alphs/Bet Spillover Curve for CL/AN
Discriminator(12)%Spill over AlphaBeta (152136

45 15

s w5 s
Discriminator Setting

165





(a)                                                                         (b)

[image: image5.png]Alpha/Beta Spillover Curve for: CL/AM
Discriminator stting (122) % Spil Over Alpha:Beta(L 66:1.28)

60

40

20 —+—Alpha
0 —i—Beta

50 100 120 140 180 130

Discriminator Setting




[image: image6.png]Alpha/Beta Spillover Curve for: CL/AM
Discriminator(121)5% Spill Over AlphaBeta;18:1.32)

1
Discriminator S¢

etting

——A
—&-B /)
0 —i L i
5 10 o oo





(c)                                                                        (d)          

[image: image7.png]‘Alpha/Beta spillover Curvefor : CL/AM
Discriminator(120)2% Spill Over AlphaiBeta 1.45:1.63)

80

——A

100 120 140 160
Discriminator Setting





                                                                                           (e)

Figures 2(a) – 2(e): Discriminator Setting 118, 124, 121,122 and 120

Table 1(a): The proportion of water, toluene, and triton-N100, appearance and the activities for  beta emitter (3H)

[image: image8.png]iatre. | Count (CP)| Efiency | Water [Toutene| Triton| Appearance | Homogenity | Viscosity s of water insoluion] Meritvalue | %Triton
W | s 05 | 8 | 1 | 1| ity | heterosenous | viscous 0 s 1
wazs | 160 15 | 7 | 2 |+ | mity | rererosenous | viscous 70 1031 1
wiso | _enn 50 | & | 3 | 1| mity | nererosenous |mor viscous s 2973 1
wisr| 72e | 77 | 5 | & | 1| miy | netsrosenous |norvisco 50 m30 1
wis2| 10 | 110 | s | 5 | 1| miy | netsrosenous |norviscous o 92 1
wiss| 730 | 82 | 3 | & | 1| miy | netsrosenous |norviscom 0 2061 1
wiss | see | 102 | 2 | 7 | 1| mily | netsrosenous | norviscou 2 2037 1
wisr| s | 00 | 1 | & | 1| miy | netsrosenous |norviscom 10 1003 1
wasi| sose | sae | 1 | 7 | 2 | miy | newsrosencus | viscow 10 3256 20
wis2 | 1748 | 184 | 2 | & | 2 | opsaue | hererogenous | viscous 2 3677 20
wass| 1920 | 143 | 3 | 5 | 2 | opsaue | hererosenous | viscous 0 a0 20
wise | o 90 | & [ & [ 2 [ opsave | neterogenous | viscous o 390 20
wass | s sé | 5 | 3 | 2| opaque | neterosenous | viscous 50 2695 20
wise| 290 | 31 | & | 2 | 2 | opsaue | hererosenous | viscous s 1533 20
was7 | 1 16 | 7 | 1 | 2 | clowsy | nererossnous | viscous 70 %61 20
wass | e 18 | & | 1 | 3| closy | nererossnous | viscous s 1110 0
wiss| e | ee | 5 | 2 | 3 | cmsr | nomossneous | notviscous 50 286 0
wigo| e | 73 | & | 3 | 3 | csr | romossneous | notviscous o 210 0
wier| 1oss | 11s | 3 | @ | 3 | cowey | nerosencus | viscow 0 225 0
wiga| 13 | 183 | 2 | 5 | 3 | cowey | newsrosencus | viscow 2 50 0
wigs | e | 27 | 1 | & | 3 | csr | nomossneous | notviscous 10 a0 0
wige | mse | 262 | 1 | 5 | & | cimsr | romossneous | norviscous 10 2017 w0
wass | 136 | 151 | 2 | 4 | 4 | ceor | nomogencour | notviscour » 020 w
wass | oo 54 | 3 | 3 | & | clouy | eterogenous | viscous 0 2013 w0
waer | sas se | & | 2 | & | cesr | romossneous | viscow o 2227 w0
wigs | 25 | 25 | 5 | 1 | & | cesr | romossneous | viscow 50 137 w0
wigs | w0 | 33 | & | 1 | 5 | cesr | romossneous | viscow o 1303 0
waro | est 66 | 3 | 2 | 5 | cesr | nomosensos | wscous 0 2054 0
wiri| ws1 | 125 | 2 | 3 | 5 | csr | nomossneous | notviscous 2 2506 0
wir2 | 10 | 196 | 1 | & | 5 | csr | nomossneous | notviscous 10 1943 0
wirs| s | 167 | 1 | 3 | 6 | cmsr | nomossneous | notviscous 10 1570 s
wire| e | ss | 2 | 2 | & | cmsr | nomossneous | notviscous 2 1850 s
wazs | _en 43 | 3 | 1 | 6 | clouy | nomossneous | norviscous 0 158 s
waze | sa1 57 | 2 | 1 | 7 | cesr | romossneous | viscow 2 1139 o
warr | 1w | 1ie | 1 | 2 | 7 | cesr | romossneous | viscow 10 185 o
wa7s | es1 72 | 1 | 1 | 8 | cear | nomoseneous | notviscous 0 716 50





        Note : The best portion of cocktail mixture 

Table 1(b): The proportion of water, toluene, and triton-N100, appearance and the activities for alpha emitter (226Ra)
[image: image9.png]vialtd _|count (cPu)| ficiency | water | Toens | Titon |dppesrance| Homogenity | Viscosity |waterinsols Meritvaive | %Titon
wazs | 203 a9 s 1 1 mitty | neterogenous | viscous | 20 36 1
waso | zse 2 7 2 1 mitty | neterogenous | viscous | 70 22 1
Wi | en s s 5 1 mitty | _neterogenous | not viscous| €0 Y 1
was2 | s 1 s s 1 mitty | _neterogenous | not viscous| 50 w053 1
wiss | e 7 s s 1 mitty | neterogenous | not viscous| 40 a2 1
wiss | sos 23 5 s 1 mitty | _neterogenous | not viscous| 30 3680 1
wass | sas 125 2 7 1 mitty | _neterogenous | not viscous| 20 2582 1
wass | s 160 T s 1 mitty | _neterogenous | not viscous| 10 1604 1
wasr | 7 6 T 7 2 mitty | neterogenous | viscous | 10 w57 20
wiss | e 156 2 s 2 opaaue | heterogenous | viscous | 20 125 20
wass | sar 102 5 s 2 opsave | heterogenous | viscous | 30 261 20
waso | e 95 s s 2 opsave | heterogenous | viscous | 40 955 20
W | o7 s s 5 2 opsave | heterogenous | viscous | 50 am3s 20
w2 | 76 s 2 2 opsave | heterogenous | viscous | 60 s 20
wass | e 95 7 T 2 clousy | neterogenous | viscous | 70 517 20
wase | s 75 s T 5 clousy | neterogenous | viscous | &0 760 0
wass | s 52 s 2 5 clear | _nomogeneous | notvscous | 50 087 0
wase | 357 7 s 3 3 clesr__|_nomogeneous | notviscous | 40 s 0
wasr | e 101 5 s 5 clousy | neterogenous | viscous | 30 3080 0
wase | e 121 2 s 5 clousy | neterogenous | viscous | 20 2011 0
wass | e 156 T s 5 clear_|_nomogeneous | notviscous | 10 1557 0
waio | s7e 160 3 s 4 clear | nomogeneous | notviscous | 10 1601 w0
waor | w5 52 2 7 7 cleor | homogeneous | notviscous | 20 1006 w
Wiz | 31 a5 3 3 s clouty | neterogenous | viscous | 30 264 w0
w0 | e o2 s 2 s clear | nomogeneous | viscous | 40 3690 w0
watos | 30 7 s T s clear | nomogeneous | viscous | 50 a1 w0
waios | a7 50 s T s clear | nomogeneous | viscous | 40 2180 0
waie | a7 ) 5 2 s clear | nomogeneous | viscous | 30 2505 0
waior | a0 ) 2 5 s clear_|_nomogeneous | notviscous | 20 1989 0
waie | s 101 T s s clear | _nomogeneous | notviscous | 10 1012 0
wais | s 23 T 5 s clear | _nomogeneous | notviscous | 10 1225 s
Wit | 3 ) 2 2 s clear_|_nomogeneous | notviscous | 20 w51 s
W | 75 5 T s clouty | nomogeneous | notviscous |30 256 s
Wz | s e 2 T 7 clear | nomogeneous | viscous | 20 1565 o
Wi | e 150 T 2 7 cear | nomogeneous | viscous | 10 1501 o
At | aa1 08 3 3 s clear | homogeneous | notviscous | | 10 w075 50





        Note : The best portion of cocktail mixture 

[image: image10.png]



Figure 2(a): Appearance 

[image: image11.png]440
420
400
380
360
340
320
300
280
260
240
220
200
180
160
140
120
100
80

650

Lo





Figure 2(b): Merit Value

[image: image12.png]



Figure 2(c): Count Rates of 3H

From Figures 2(a), (b) and (c) are the contour for appearance, merit value and count rate of 3H for 36 vials. It gives clear indication of the best appearance, merit value and count rate. 

Conclusion

This instrument, LSC Packard TRICAB 2700 used 14C for its calibration. The best discriminator setting for alpha/beta counting obtained in this study is 120 with 1.45:1.63 percentage ratio of alpha/beta spillover. The best proportion for the scintillation cocktail is found to be 2:4:4 for water, toluene, and triton N-100 with 4 g/L PPO and 0.4 g/L POPOP. 

Acknowledgement

The authors would like to thank Malaysian Nuclear Agency for allowing the project to be carried out at Malaysian Nuclear Agency (NM) laboratory in Bangi.
References

1. Packard, (1994) Liquid Scintillation Counter (TRICAB 2700) Manual.
2. UNSCEAR (1993) Ionizing Radiation: Sources and Effects of Ionizing Radiation. United Nations Scientific Committee on the Effects of Atomic Radiation. Report to the General Assembly, with Scientific Annexes. United Nations Sales Publication E.94.IX.2.United Nations, New York.
3. WHO, (1993). Guidelines for Drinking Water Quality. Second ed., Recommendations, vol. 1, World Health Organization, Geneva, Switzerland, pp. 114–121.
4. UNSCEAR (United Nations Scientific Committee on the Effects of Atomic Radiations), (2000). Sources and effects of ionizing radiation. Report to the General Assembly, United Nations, New York.
5. Malanca, A., Repetti, M., Macedo, H.R., (1998). Gross alpha and beta-activities in surface and ground water of Rio Grando do Norte, Brazil. Inter J. Appl. Radiat.and  Isotopes. 49 (7), 893–898.
6. Terlikowska T, D Hainos, P Cassette  and T Radoszewski (2000). ‘Application of alpha/beta discrimination in liquid scintillation counting for the purity control of 99mTc medical solutions’. Appl. Radiat and Isotopes, 52, 627-632.
7. Shing-Fa Fang et al (2006). ‘Comparison of Alpha/Beta Separation Performance of Commercially Available Scintillation Cocktails Counting by QuantulusTM 1220 Liquid Scintillation Counter’ Division of Health Physics, Institute of Nuclear Energy Research, 1000 Wen-Hua Road, Chiaan Village, Lungtan, Taoyuan, Taiwan, R.O.C.
8. Maurizio Forte , Rosella Rusconi, Elisabetta Di Caprio, Silvia Bellinzona and Giuseppe (2000) ‘Natural Radionuclides Measurements in Drinking Water by Liquid Scintillation Counting’. ARPA Loambardia, Italy. 
9. Cantaloub, M. J. Higginbothram, J. Istok and L semprini. (2008) ‘Interaction of sample, cocktail and headspace volume when measuring aqueous Radon in small volume. Nuclear Engineering Dept., Oregon State University. 
10. Van Der Laarse, J.D, (1967) Experience with emulsion counting of tritium, Inter.J.Appl.Radiat. Isotopes, 18, 485.
Organic Solvent





3.0 ml











36
27

_1352716279.xls
Chart1

		85		85

		97		97

		109		109

		121		121

		133		133

		145		145

		157		157



Alpha

Beta

Discriminator Setting

%  SpillOver

Alpha/Beta Spillover Curve for CL/AM
Discriminator(124);%Spill over Alpha:Beta  (1.52:1.36)

0.12

33.03

0.18

14.27

0.48

4.06

1.67

1.31

4.98

0.48

11.88

0.28

22.97

0.14



Set 1

		Beta A		Beta B		Alpha		CPMA		CPMB		CPMC		SIS				Beta A		Beta B		Alpha		CPMA		CPMB		CPMC		SIS

		0-500		0-40		300-500		1724		444		1		413.32				0-500		0-40		300-500		16018		5280		19		359.15

		0-1000		0-20		100-200		1827		227		11		528.89				0-1000		0-20		100-200		16711		2478		85		414.76

		0-100		0-10		0-50		955		9		2		149.88				0-100		0-10		0-50		9562		942		33		138.72

		0-100		0-50		50-100		981		544		6		149.48				0-100		0-50		50-100		9463		5815		71		140.73

		0-100		0-20		100-200		1025		239		21		147.65				0-100		0-20		100-200		9425		2157		109		141.17

		0-100		0-15		200-300		942		164		1		146.26				0-100		0-15		200-300		9465		1513		23		143.24

		0-100		0-16		300-400		954		186		3		144.11				0-100		0-16		300-400		9421		1597		11		145.2

		0-100		0-17		400-500		952		196		0		141.95				0-100		0-17		400-500		9430		1641		3		146.78

		0-100		0-18		500-600		928		18		1		146.16				0-100		0-18		500-600		9277		1746		6		147.58

		0-100		0-19		600-700		992		240		0		144.49				0-100		0-19		600-700		9220		1904		29		144.92

		0-100		0-14		700-800		972		163		0		141.9				0-100		0-14		700-800		9275		1237		33		147.31

		0-100		0-13		800-900		954		137		0		147.94				0-100		0-13		800-900		9194		1129		15		148.41

		0-100		0-12		900-1000		936		139		0		144.56				0-100		0-12		900-1000		9120		1040		2		148.94

		0-100		0-11		1000-1100		908		104		0		146.86				0-100		0-11		1000-1100		9345		881		0		150.54

		0-100		0-8		1200-1300		942		90		0		142.78				0-100		0-8		1200-1300		8992		566		0		151.56

		0-100		0-7		130-1400		972		67		0		147.58				0-100		0-7		130-1400		8988		464		0		151.91

		0-100		0-5		0-900		949		44		28		147.39				0-100		0-5		0-900		8734		253		302		151.71

		0-100		0-6		0-1000		913		64		24		141.2				0-100		0-6		0-1000		8798		410		315		152

		0-100		0-4		0-800		974		50		31		141.34				0-100		0-4		0-800		8544		201		331		154

		0-100		0-3		0-800		940		20		28		142.64				0-100		0-3		0-800		8583		1263		322		153.54

		100-200		0-18		700-800		374		197		0		230.77				100-20		0-18		700-800		3656		1499		28		245.44

		200-300		0-18		600-700		142		183		0		287.61				200-300		0-18		600-700		1718		1482		8		314.77

		300-400		0-18		500-600		147		207		0		360.41				300-400		0-18		500-600		1452		1552		8		392.91

		400-500		0-18		400-500		98		20		0		413.76				400-500		0-18		400-500		1177		1482		3		471.25

		500-600		0-18		0-900		55		218		23		466.14				500-600		0-18		0-900		813		1525		322		526.98

		600-700		0-18		500-900		44		209		0		505.74				600-700		0-18		500-900		469		1509		22		569.92

		300-800		5-20		300-600		330		186		0		512.22				300-800		5-20		300-600		4044		1430		15		580.27

		100-500		0-50		100-500		727		557		12		420.3				100-500		0-50		100-500		8129		4384		248		466.61

		0-500		0-50		0-300		1717		531		22		420.74				0-500		0-50		0-300		16463		4378		268		471.77

		0-0		0-0		0-800		0		0		27		0				0-0		0-0		0-800		0		0		300		0

		0-0		0-0		0-2000		0		0		18		0				0-0		0-0		0-2000		0		0		347		0

		0-0		0-0		0-1500		0		0		32		0				0-0		0-0		0-1500		0		0		333		0

		0-0		0-0		0-10		0		0		0		0				0-0		0-0		0-10		0		0		0		0

		0-0		0-0		0-20		0		0		0		0				0-0		0-0		0-20		0		0		0		0

		0-0		0-0		20-50		0		0		1		0				0-0		0-0		20-50		0		0		16		0

		0-0		0-0		30-300		0		0		23		0				0-0		0-0		30-300		0		0		252		0

		0-0		0-0		30-500		0		0		33		0				0-0		0-0		30-500		0		0		293		0

		0-100		0-20		100-200		945		223		15		147.17				0-100		0-20		100-200		8256		1735		185		155.75

		0-1500		0-5		0		1848		48		0		535.4				0-1500		0-5		0		17625		283		0		580.28

		0-2000		0-18		0		1856		184		0		551.62				0-2000		0-18		0		17638		1511		0		577.07

		0-1200		0-10		0		1779		127		0		502.75				0-1200		0-10		0		17345		675		0		584.22

		0-1000		0-15		0		1758		166		0		540.2				0-1000		0-15		0		17554		1283		0		575.31

		0-900		0-20		0		1786		222		0		517.76				0-900		0-20		0		17544		1694		0		575.78

		0-800		0-18		0		1857		189		0		518.54				0-800		0-18		0		17649		1475		0		576.68

		0-700		0-18		0		1818		192		0		520.09				0-700		0-18		0		17491		1514		0		557.98

		0-600		0-18		0		1729		209		0		453.14				0-600		0-18		0		17101		1531		0		529.11

		0-500		0-18		0		1727		209		0		413.43				0-500		0-18		0		16260		1479		0		472.68

		0-400		0-18		0		1557		203		0		368.41				0-400		0-18		0		15177		1471		0		398.47

		0-100		0-15		200-300		1456		183		0		286.94				0-100		0-15		200-300		13400		1503		0		316

		0-200		0-18		0		1333		232		0		226.79				0-200		0-18		0		12047		1514		248.81		0

		0-300		0-18		0		1437		191		0		300.83				0-300		0-18		0		13586		1464		0		318.73

		0-1100		0-18		0		1868		222		0		508.81				0-1100		0-18		0		17416		1498		0		574.89

		0-200		0-0		0-100		1254		0		8		223.6				0-200		0-0		0-100		13344		0		91		215.08

		0-122		0-18		0-122		1068		183		11		171.14				0-122		0-18		0-122		1068		183		11		171.14

		0-0		0-0		0-122		0		0		9

		0-2000		0-0		0-2000		1793		0		25						0-2000		0-0		0-2000		17196		0		303

		80-182		0-20		80-182		501		225		27		222.58				80-182		0-20		80-182		4285		2234		100		214.27

		Set 1: Menggunakan Hisafe ALPHA/BETA

		Spike 0.1 ml into 20 ml Hisafe (A); Aperance; Solution light

		Spike 1 ml into 19 ml Hisafe (B)						; Appearance; Cloudy





Set 2

		Beta A		Beta B		Alpha				Time:1 minit

		0-2000		0-0		0-1000

		CPMA		CPMB		CPMC		SIS		Appearance

		1751		208		16		547.59		Light

		17560		1664		368		584.76		Cloudy

		CPMA		CPMB		CPMC		SIS		Appearance				Sr 90		57.744 Bq/ml

		56		23		12		284.57

		42		6		5		441.87

		54		18		6		310.24

		28		4		10		433.75

		43		17		3		240.88

		46		20		7		165.59

		54		16		12		499.36

		37		10		2		447.34

		54		12		5		223.15

		45		20		4		265.88

		43		14		10		616.34

		44		13		9		591.3

		44		10		1		491.42

		51		13		7		252.85

		65		23		2		212.76

		42		11		14		700.93

		40		8		12		683.17

		47		11		6		56.99

		55		14		1		465.26

		47		11		6		234.35

		49		16		1		197.63

		42		11		7		518.76

		50		11		7		689.7

		47		12		3		519.63

		64		16		4		397.06

		57		16		5		547.46

		40		20		4		452.61

		52		6		6		187.08

		50		20		5		719.07

		56		10		8		524.61

		48		12		3		704.43

		47		9		8		391.84

		59		12		3		390.57

		47		8		6		671.62

		46		3		5		309.35

		60		20		8		271.06





set 3

		Beta A		Beta B		Alpha				Time:2 minit														Beta A		Beta B		Alpha				Time:2 minit

		0-2000		0-0		0-1000																		0-2000		0-0		0-1000

		CPMA		CPMB		CPMC		SIS		Appearance														CPMA		CPMB		CPMC		SIS		Appearance

		1817		225		29		543.6		Light				0.1 ml Sr90+ 0.1 ml P0209 + 19 ml coktail Hisafe								1		2366.5		1228		23.5		105.77						Sr 90		57.744 dpm/g

		17560		1662		359.5		594.11		Cloudy				1 ml Sr90+ 1 ml P0209 + 18 ml coktail Hisafe								2		2548.5		687.5		24		225.91

		CPMA		CPMB		CPMC		SIS		Appearance				PO 209		33.5730 dpm/g						3		245.5		1163		11.5		123.01

		54		22		6.5		186.36														4		2655		624.5		11		248.55

		41		10		7.5		325.33														5		2652		1029.5		13.5		155.81

		52		26		7.5		162.18														6		2507.5		1257.5		35.5		113.26

		38		8		6		565.51														7		2532.5		466		23		330.09

		46		15.5		5.5		231.17														8		2669		734.5		12		212.96

		57.5		24.5		7.5		167.29														9		2604.5		1124.5		7		136.01

		42.5		8.5		9.5		634.04														10		2476		1370		4		96.947

		44.5		15.5		7.5		468.81														11		2575		415		21.5		395.39

		39		13.5		10.5		312.57														12		2649.5		466.5		9.5		333.9

		43.5		15		8.5		225.8														13		2535.5		865.5		16		174.9

		35		4.5		7		644.77														14		2638		953		34.5		172.49

		47.5		11.5		9.5		447.3														15		2585		1056		32.5		145.75

		48		8.5		1		389.43														16		2624.5		319.5		18.5		476.69

		52.5		13.5		6		287.53														17		2641		477.5		13		307.5

		49		18		4		224.45														18		2607.5		535		31.5		275.37

		36.5		8		12		726.26														19		2856.5		808		7.5		222.82

		41.5		9		7.5		696.55														20		2605		710		30		223.05

		51		11.5		5		433														21		2338		1521.5		5.5		68.699

		45		8.5		6.5		366.67														22		2401		896.5		39.5		292.31

		56.5		14		9.5		237.22														23		2631		365.5		12.5		449.72

		41		14.5		2		153.73														24		2575		565		10.5		301.71

		50		11.5		12		1044.7														25		2687		563.5		24		289.91

		44.5		7		10.5		735.35														26		2668		727		6.5		214.32

		39.5		7.5		4.5		648.19														27		2624.5		981		4		161.16

		43.5		8		6		535.27														28		2449		1411.5		6		88.375

		39.5		9.5		4.5		591.44														29		2548.5		743		24.5		265.82

		46.5		12		6.5		300.51														30		2658		323		34		466.13

		43.5		12.5		5		246.64														31		2736		513		12		301.95

		45.5		8		7		678.71														32		2693.5		661		25		235.21

		42		12		7		789.74														33		2686.5		480		5.5		324.55

		53.5		11.5		8		772.81														34		2686.5		553		31		278.18

		45.5		12		8		561.42														35		2737		492		21		322.02

		57		13.5		5		631.91														36		2476		1222		5		113.08

		48		14.5		4.5		450.54

		60		15		2.5		280.69

		54		15.5		3		321.58





Sheet1

		Setting		Alpha		Beta																								%Spillover

		90		0.16		23.55																				Setting in use:118				Alpha		Beta

		102		0.37		7.7																				Computed setting				1.47		1.5

		114		0.97		2.29																								1.47		1.5

		126		3.26		0.74

		138		8.64		0.28

		150		17.57		0.11

		162		30.79		0.08

		Setting		Alpha		Beta

		85		0.12		33.03

		97		0.18		14.27																								%Spillover

		109		0.48		4.06																				Setting in use:120				Alpha		Beta

		121		1.67		1.31																				Computed setting				1.52		1.36

		133		4.98		0.48																								1.52		1.36

		145		11.88		0.28

		157		22.97		0.14

		Setting		Alpha		Beta

		80		0.16		44.41																						Setting in use:118				Alpha		Beta

		97		0.22		13																						Computed setting				1.47		1.43

		114		1.05		2.32																										1.47		1.43

		131		4.93		0.53

		148		16.61		0.15

		165		34.72		0.11

		182		53.63		0.08

																																Alpha		Beta

		Setting		Alpha		Beta																						Setting in use:122				1.66		1.28

		80		0.13		50.28																						Computed setting: 122				1.47		1.43

		95		0.18		20.62

		110		0.57		4.86

		125		2.21		1.13

		140		8.21		0.31

		155		19.85		0.12

		170		36.96		0.09

																																Alpha		Beta

		Setting		Alpha		Beta																						Setting in use:122				1.66		1.28

		80		0.12		48.52																						Computed setting: 122				1.47		1.43

		96		0.17		18.4

		112		0.7		3.92																						80-175

		128		3.41		0.82

		144		10.93		0.29

		160		26.2		0.12

		176		44.72		0.09

		Setting		Alpha		Beta

		80		0.15		49.35																										Alpha		Beta

		97		0.19		17.27																						Setting in use:122				1.61		1.27

		114		0.74		3.41																						Computed setting: 122				1.61		1.27

		131		3.85		0.64

		148		13.86		0.18

		165		31.28		0.1

		182		50.68		0.08

																																Alpha		Beta

		Setting		Alpha		Beta																						Setting in use:120				1.45		1.63

		85		0.14		37.36																						Computed setting: 120				1.45		1.63

		102		0.28		10.02

		119		1.31		1.83

		136		6.58		0.42

		153		19.26		0.14

		170		38.37		0.08

		187		57.73		0.06

																																Alpha		Beta

																												Setting in use:120				1.45		1.63

		Setting		Alpha		Beta																						Computed setting: 120				1.45		1.63

		70		0.13		63.69

		85		0.15		37.77

		100		0.31		11.88

		115		1		2.75

		130		4.12		0.67

		145		12.37		0.23

		160		26.63		0.13

		Setting		Alpha		Beta

		90		0.16		26.9																										Alpha		Beta

		102		0.31		9.7																						Setting in use:120				1.51		1.56

		114		0.82		2.96																						Computed setting: 120				1.51		1.56

		126		2.74		0.93

		138		7.53		0.35

		150		16.18		0.16

		162		28.87		0.13

		Setting		Alpha		Beta																										Alpha		Beta

		95		0.2		18.25																						Setting in use:120				1.55		1.63

		106		0.42		6.59																						Computed setting: 120				1.55		1.63

		117		1.17		2.22

		128		3.33		0.77

		139		8.36		0.29

		150		16.25		0.19

		161		27.54		0.12

		Setting		Alpha		Beta

		110		0.59		4.9																										Alpha		Beta

		122		1.74		1.38																						Setting in use:121				1.55		1.57

		134		5.22		0.53																						Computed setting: 121				1.55		1.57

		146		12.14		0.22

		158		23.13		0.12

		170		36.92		0.08

		182		50.97		0.06

		Setting		Alpha		Beta																										Alpha		Beta

		80		0.12		48.86																						Setting in use:121				1.57		1.5

		93		0.18		23.43																						Computed setting: 121				1.57		1.5

		106		0.37		7.26

		119		1.28		1.87

		132		4.42		0.56

		145		11.19		0.21

		158		22.99		0.13





Sheet1

		



Alpha

Beta

Discriminator Setting

% SpillOver

Alpha/Beta Spillover Curve for:Cl/Am
Discriminator(118);% Spill Over Alpha:Beta (1.47:1.5)



		



Alpha

Beta

Discriminator Setting

%  SpillOver

Alpha/Beta Spillover Curve for CL/AM
Discriminator(120);%Spill over Alpha:Beta  (1.52:1.36)



		



Alpha

Beta

Discriminator Setting

%  Spill Over

Alpha/Beta Spillover Curve for:CL/AM



		



Alpha

Beta

Discriminator Setting

% Spill Over

Alpha/Beta Spillover Curve for: CL/AM



		



Alpha

Beta

Discriminator Setting

% Spill Over

Alpha/Beta Spillover Curve for: CL/AM



		



Alpha

Beta

Discriminator Setting

% Spill Over

Alpha/Beta Spillover Curve for: CL/AM



		



Alpha

Beta

Discriminator Setting

% Spill Over

Alpha/Beta SpillOver Curve for : CL/AM



		



Alpha

Beta

Discriminator Setting

% Spill Over

Alpha/Beta Spillover Curve for : CL/AM



		



Alpha

Beta

Discriminator Setting

% SpillOver

Alpha/Beta Spillover Curve for: CL/AM



		



Alpha

Beta

Discriminator Setting

% Spill Over

Alpha/Beta SpillOver Curve for: CL/AM



		



Alpha

Beta

Discriminator Setting

% SillOver

Alpha/Beta SpillOver Curve for : CL/AM



		



Alpha

Beta

Discriminator Settinge

% SpillOver

Alpha Beta SpillOver Curve for: Cl/Am




