The Malaysian Journal of Analytical Sciences Vol 16 No 1 (2012): 71 – 78

DETERMINATION OF VITAMIN E ISOMERS IN PLASMA USING ULTRA PERFORMANCE LIQUID CHROMATOGRAPHY
(Penentuan Isomer-Isomer Vitamin E di dalam Plasma Menggunakan Kromatografi Cecair Berprestasi Lampau)

Shahidee Zainal Abidin2,3, Mohammad Johari Ibahim2, Mohd Hamim Rajikin1 

and Nuraliza Abdul Satar1(
1Department of Physiology, 

2Institute of Medical Molecular Biotechnology,

3Centre for Pathology Diagnostic and Research Laboratories,

 Faculty of Medicine, 

Universiti Teknologi MARA, 40450 Shah Alam, Selangor, Malaysia
*Corresponding author: nuraliza064@salam.uitm.edu.my
Abstract

A rapid method for quantification of vitamin E isomers (α-, β-, γ-, δ- tocopherol and tocotrienols) was developed by using ultra performance liquid chromatography (UPLC) from human plasma. This method involved liquid-liquid extraction of plasma sample to extract the lipophilic portion for injection into UPLC system. The UPLC system was equipped with fluorescent detector (296 nm excitation, 330 nm emissions) and C18 column with 1.7 µm particle size (2.1 x 110 mm column) for particle separation. The mobile phase comprises of 1.0% of H2O and 99.0% of methanol (HPLC grade) with flow rate adjusted at 0.4 ml min-1. Peak separations were accomplished after 3.50 minutes and 5 µl treated samples were required for injection. Results obtained showed that the UPLC system was able to separate, detect and quantitate six peaks, namely α-, γ-, δ- tocopherol and tocotrienols, respectively. However, it failed to separate and quantificate both, β- tocopherol and tocotrienol due to low amount in nature. The calibration curve was linear for a concentration ranging from 1 µg ml-1 – 10 µg ml-1 of total vitamin E standard which r2 > 0.9996 for all isomers. The within day and between day coefficients of variation were less than 10.29% (n=4) and 10.73% (n=4) for all isomers. The accuracy of isomers quantification for 1 µg ml-1 (n=4) and 10 µg/ml (n=4) vitamin E stocks were 93.52% - 98.13 %. The within day and between day coefficients of variation for vitamin E standard were less than 7.59% (n=4) and 9.74% (n=4) for all isomers while for plasma sample were less than 10.29% (n=4) and 10.73% (n=4) for all isomers. The recovery rates for all isomer were ranging from 98.96 ± 7.22 and 107.50 ±1.19 and the limit of detection was 2.53 ng. The present data suggested that this method required extremely short analysis time i.e. 2.50 min, required less samples for injection and excellent chromatographic reproducibility.

Keywords:  UPLC, Tocotrienols, Tocopherol, Plasma, Vitamin E
References
1. Qureshi AA, Karpen CW, Nor RM, Qureshi N, Papasian CJ, Morrison DC and Folts JD (2011). Tocotrienol-Induced Inhibition of Platelet Thrombus Formation and Platelet Aggretation in Stenosed Canine Coronary Arteries. Lipid in Health and Disease, 10:58, 1-13.
2. Tan B (2005). Appropriate Spectrum Vitamin E and New Perspectives on Desmethyl Tocopherols and Tocotrienols. The Journal of the American Nutraceutical Association, 8:1, 35-42.
3. Llacuna L, Fernández A, Von Montfort C, Matías N, Martínez L, Caballero F, Antoni Rimola, Montserrat Elena, Morales A, Fernández-Checa JC and García-Ruiz C (2011). Targeting Cholesterol at Different Levels in the Mevalonate Pathway Protects Fatty Liver against Ischemia–Reperfusion Injury. Journal of Hepatology, 54, 1002–1010
4. Pearce BC, Parker RA and Deason ME (1992). Hypocholesterolemic Activity of Synthetic and Natural Tocotrienols. Journal of Medicine Chemistry, 35:20, 3595–3606.
5. Schauss AG (2009). Tocotrienols: A Review. In Watson RR and Preedy VR, Tocotrienols Vitamin E Beyond Tocopherols. CRC Press Taylor & Francis Group. 3-12.
6. Abidi S. (2000). Chromatographic Analysis of Tocol-Derived Lipid Antioxidants. Journal of Chromatography A, 881:1-2, 197-216.
7. Citová I, Havlíková L, Urbánek L, Solichová D, Nováková L and Solich P (2007). Comparison of a Novel Ultra-Performance Liquid Chromatographic Method for Determination of Retinol and α-Tocopherol in Human Serum with Conventional HPLC using Monolithic and Particulate Columns. Anal Bioanal Chem,  388:3, 675–681.
8. Paliakov EM, Crow BS, Bishop MJ, Norton D, George J and Bralley JA  (2009). Rapid Quantitative Determination of Fat-Soluble Vitamins and Coenzyme Q-10 in Human Serum by Reversed Phase Ultra-High Pressure Liquid Chromatography with UV Detection. Journal of Chromatography B, 877:1-2, 89-94.
9. Xu Z (2002). Analysis of Tocopherols and Tocotrienols. In Wrolstad RE, Current Protocols in Food Analytical Chemistry. John Wiley & Sons Inc. D1.5.1-D1.5.12.
10. Fountain KJ, Grumbach ES, Xu J, Messina CJ and Diehl MD (2008). Application of Novel Ethylene Bridged Hybrid Particles for Hydrophilic Interaction Chromatography. Milford: Water Corporation.
