The Malaysian Journal of Analytical Sciences, Vol 16 No 2 (2012): 142 - 147
Samat, S. B. et al:   
STABILITY OF THE ABSORBED DOSE TO WATER CALIBRATION COEFFICIENT, ND,w FOR THERAPY LEVEL IONIZATION CHAMBERS BELONGING TO LOCAL RADIOTHERAPY CENTRES: ANALYSIS OF RESULTS OBTAINED DURING 2004 -2010


[image: image1.png]



STABILITY OF THE ABSORBED DOSE TO WATER CALIBRATION COEFFICIENT, ND,w FOR THERAPY LEVEL IONIZATION CHAMBERS BELONGING TO LOCAL RADIOTHERAPY CENTRES: ANALYSIS OF RESULTS OBTAINED DURING 2004 -2010

(Kestabilan Pekali Kalibrasi Dos Terserap Terhadap Air, ND,w Untuk Kebuk Pengionan Tahap Terapi Kepunyaan Pusat Radioterapi Tempatan: Analisis Keputusan 

Yang Diperoleh Sepanjang 2004-2010)
S. B. Samat1*, W. Priharti1, S. S. Chong1, T. Kadni2 and M.T. Dolah2
1School of Applied Physics,

Faculty of Science and Technology, Universiti Kebangsaan Malaysia
43600 UKM Bangi, Selangor

2Secondary Standard Dosimetry Laboratory (SSDL)

Malaysian Nuclear Agency (Nuclear Malaysia) 

43000 Kajang, Selangor
*Corresponding author: sbsamat@ukm.my
Abstract

According to the IAEA, the calibration of therapy level ionization chambers in terms of absorbed dose to water calibration coefficient, ND,w must be within ± 1.5% acceptance limit. This is for the purpose of getting accurate absorbed dose for the patient undergoing radiotherapy treatment. The objective of this work is to evaluate the deviation of ND,w for 29 chambers belonging to 16 local radiotherapy centres. Eight types of chambers have been calibrated at the SSDL Laboratory of Malaysian Nuclear Agency for the period of seven years. The mean µ of the ND,w deviation together with its standard error (SE) and standard deviation σN-1 were calculated. It is found that out of 29 chambers, 26 chambers yielded µ ± SE within the permitted value of ± 1.5%. For the other three chambers, despite their values of µ ± SE lie slightly outside the range of          ± 1.5%, they are still within the 95% confident interval of the ± 1.5%. It is concluded that the ND,w of the chambers belonging to the local radiotherapy chambers are stable in their performance the dose to the patients measurement.

Keywords:   absorbed dose to water calibration coefficient ND,w, local radiotherapy centres, percentage deviation, SSDL Malaysia, stability of ND,w
Abstrak

Menurut IAEA, julat penerimaan pekali kalibrasi dos terserap terhadap air, ND,w untuk kebuk pengionan tahap terapi mesti berada antara ± 1.5%. Hal ini bertujuan untuk mendapatkan dos terserap yang tepat bagi pesakit yang menjalani rawatan radioterapi. Objektif kerja ini adalah untuk menilai sisihan ND,w bagi 29 kebuk pengionan kepunyaan 16 pusat radioterapi tempatan. Lapan jenis kebuk telah dikalibrasi di makmal SSDL Agensi Nuklear Malaysia dalam tempoh tujuh tahun. Nilai purata µ untuk sisihan ND,w bersama dengan nilai ketidakpastian standard (SE) dan juga sisihan standard σN-1 dikira. Didapati bahawa 26 daripada 29 kebuk menunjukkan nilai µ ± SE berada dalam julat yang dibenarkan iaitu ± 1.5%. Walaupun nilai    µ ± SE untuk tiga kebuk yang lain berada di luar julat ± 1.5%, namun ianya masih  berada dalam aras keyakinan 95% daripada ± 1.5%. Dapat disimpulkan bahawa nilai ND,w untuk kebuk pengionan kepunyaan pusat radioterapi tempatan adalah stabil bagi pengukuran dos kepada pesakit.

Kata kunci: Pekali kalibrasi dos terserap terhadap air, ND,w, pusat radioterapi tempatan, peratus sisihan, SSDL Malaysia, kestabilan ND,w
Introduction

It has been well documented that accuracy in the beam calibration of the radiotherapy machines would contribute to the accurate dose delivery to the patients [1,2]. Therapy level ionization chamber (IC) is the important tool for the calibration, therefore its accurate absorbed dose to water calibration coefficient, ND,w certainly would contribute to the accurate dose delivery [3]. It is anticipated that the ND,w of an IC would change with time, therefore the IAEA has recommended that an IC should be calibrated every year for its ND,w [4]. 

For the determination of ND,w, SSDL Malaysia of Malaysian Nuclear Agency (Nuclear Malaysia) has participated in the postal IC intercomparison programmes [5]. Some of the results are shown in Table 1, which shows that the laboratory has provided therapy-level ND,w values within the internationally accepted standards i.e. within the ± 1.5%. For this reason, Nuclear Malaysia is capable for providing the ND,w values for IC belonging to the local radiotherapy centres (RC).

The purpose of this work is to study the ND,w values of 29 ICs belonging to the 16 RC, from the year 2004 to 2010. These ND,w values were determined by Nuclear Malaysia. The objective is to find out how the ND,w values vary with time i.e. its stability.

Table 1.
The intercomparison results of ND,w between Nuclear Malaysia (NM) and IAEA

	Calibration Year
	IC Model
	ND,w (mGy/nC)
	Deviation 
(%)

	
	
	NM
	IAEA
	

	1995
	NE 2581
	57.20
	56.79
	  0.72

	1997
	NE 2581
	57.08
	57.07
	  0.02

	1998
	NE 2571
	45.33
	45.39
	- 0.13

	2002
	M 30001
	52.46
	52.35
	   0.21

	2004
	M 30001
	52.52
	52.35
	   0.32


Experimental
Figure 1 shows the experimental set-up for the determination of  ND,w. The method follows the procedures of substitution method recommended by the IAEA [6]. The PMMA water phantom is of size 30 cm × 30 cm × 30 cm. A waterproof sleeve is used to place the IC at its reference point at a depth 5 g/cm2 in the water phantom. In this work a reference standard chamber of 0.6 cm3 NE2571(#1028) is used. 

Table 2 lists the type of the 29 ICs belonging to the 16 RCs. The table also shows the ND,w calibration frequency for a period of seven years from the year 2004 to 2010. The IAEA requires each IC to be calibrated every year for purpose of getting the most recent value of ND,w. However in this work, the determination of RC’s ND,w is based on the request by the RC themselves not the request by the SSDL Malaysia. Whenever SSDL Malaysia receives the IC, the calibration work will be done immediately. In other words, the frequency of each IC calibration is determined by the radiotherapy centres themselves. 

[image: image2.jpg]Electrometer connector

%Co teletherapy machine Thermometer

Ton chamber holder

Ton chamber

- FS =10 cm x 10 cm

Perspex casing

Water phantom
Calibration table

SSD =100 cm

SCD =105 cm





Figure 1. Experimental setup of ND,w calibration
Table 2.   ND,w calibration frequency for chambers belonging to local radiotherapy centres

	IC Model
	IC No
	RC No
	Freq.
	ND,w calibration frequency 

	
	
	
	
	2004
	2005
	2006
	2007
	2008
	2009
	2010

	CC 13#7063
	1
	4
	3
	
	
	
	[image: image3.png]



	[image: image4.png]



	[image: image5.png]



	

	FC 65-G#1328
	2
	3
	3
	
	
	
	[image: image6.png]



	[image: image7.png]



	[image: image8.png]



	

	FC 65-G#1371
	3
	5
	4
	[image: image9.png]



	[image: image10.png]



	[image: image11.png]



	[image: image12.png]



	
	
	

	FC 65-G#367
	4
	7
	7
	[image: image13.png]



	[image: image14.png]



	[image: image15.png]



	[image: image16.png]



	[image: image17.png]



	[image: image18.png]



	[image: image19.png]




	FC 65-G#368
	5
	4
	6
	[image: image20.png]



	
	[image: image21.png]



	[image: image22.png]



	[image: image23.png]



	[image: image24.png]



	[image: image25.png]




	FC 65-G#369
	6
	4
	4
	
	
	[image: image26.png]



	[image: image27.png]



	[image: image28.png]



	[image: image29.png]



	

	FC 65-G#371
	7
	5
	3
	
	
	
	
	[image: image30.png]



	[image: image31.png]



	[image: image32.png]




	FC 65-G#744
	8
	14
	6
	[image: image33.png]



	
	[image: image34.png]



	[image: image35.png]



	[image: image36.png]



	[image: image37.png]



	[image: image38.png]




	FC 65-G#WD 362
	9
	11
	3
	[image: image39.png]



	[image: image40.png]



	
	[image: image41.png]



	
	
	

	FC 65-G#WD 445
	10
	13
	6
	[image: image42.png]



	[image: image43.png]



	[image: image44.png]



	[image: image45.png]



	[image: image46.png]



	[image: image47.png]



	

	FC 65-G#WD 878
	11
	11
	5
	
	[image: image48.png]



	[image: image49.png]



	[image: image50.png]



	[image: image51.png]



	[image: image52.png]



	

	FC 65-G#WD1219
	12
	6
	5
	
	
	[image: image53.png]



	[image: image54.png]



	[image: image55.png]



	[image: image56.png]



	[image: image57.png]




	FC 65-P#1437
	13
	8
	3
	
	
	
	[image: image58.png]



	[image: image59.png]



	
	[image: image60.png]




	FC 65-P#742
	14
	14
	5
	[image: image61.png]



	
	[image: image62.png]



	
	[image: image63.png]



	[image: image64.png]



	[image: image65.png]




	NE 2571#2257
	15
	8
	4
	[image: image66.png]



	[image: image67.png]



	[image: image68.png]



	[image: image69.png]



	
	
	

	NE 2571#2798
	16
	15
	5
	[image: image70.png]



	
	[image: image71.png]



	[image: image72.png]



	[image: image73.png]



	[image: image74.png]



	

	NE 2581#1180
	17
	12
	6
	
	[image: image75.png]



	[image: image76.png]



	[image: image77.png]



	[image: image78.png]



	[image: image79.png]



	[image: image80.png]




	NE 2581#621
	18
	3
	7
	[image: image81.png]



	[image: image82.png]



	[image: image83.png]



	[image: image84.png]



	[image: image85.png]



	[image: image86.png]



	[image: image87.png]




	PPC 40#310
	19
	13
	4
	
	
	[image: image88.png]



	[image: image89.png]



	[image: image90.png]



	[image: image91.png]



	

	PPC 40#314
	20
	9
	3
	
	
	
	[image: image92.png]



	[image: image93.png]



	[image: image94.png]



	

	TM 30013#1745
	21
	10
	4
	
	
	[image: image95.png]



	[image: image96.png]



	[image: image97.png]



	[image: image98.png]



	

	TM 31010#0426
	22
	5
	4
	
	
	[image: image99.png]



	[image: image100.png]



	
	[image: image101.png]



	[image: image102.png]




	TW 30001#2120
	23
	15
	5
	
	
	[image: image103.png]



	[image: image104.png]



	[image: image105.png]



	[image: image106.png]



	[image: image107.png]




	TW 30013#1646
	24
	1
	4
	
	
	[image: image108.png]



	[image: image109.png]



	[image: image110.png]



	[image: image111.png]



	

	TW 30013#1681
	25
	16
	5
	
	
	[image: image112.png]



	[image: image113.png]



	[image: image114.png]



	[image: image115.png]



	[image: image116.png]




	TW 30013#2139
	26
	16
	4
	
	
	[image: image117.png]



	[image: image118.png]



	[image: image119.png]



	[image: image120.png]



	

	TW 31010#1005
	27
	16
	4
	
	
	[image: image121.png]



	[image: image122.png]



	[image: image123.png]



	[image: image124.png]



	

	W 30001#1210
	28
	15
	5
	[image: image125.png]



	[image: image126.png]



	
	[image: image127.png]



	[image: image128.png]



	[image: image129.png]



	

	WDIC 70#062
	29
	2
	5
	[image: image130.png]



	[image: image131.png]



	 
	[image: image132.png]



	[image: image133.png]



	[image: image134.png]



	 


Table 3 shows an example of ND,w calibration results for IC=4 and RC=7. Since there are seven values of ND,w, there are only six deviations ((%) that can be calculated, as the first ND,w was taken as the standard. These six deviations were then used to get the mean µ, standard error SE and standard deviation σN-1. The SE and σN-1 are calculated at 68.26% confident interval (CI). The same calculation was repeated for the other ICs. 
In this work, the results of µ ± SE and µ ± σN-1 were compared with the deviation limit of ± 1.5% accepted by the IAEA [7]. If µ ± SE is within the limit, it is said that the chamber is stable in maintaining its ND,w.
Table 3. The results of ND,w for chamber FC 65-G#367 calibrated each year from 2004-2010

	IC No
	RC No 
	Calibration Year
	ND,w (mGy/nC)
	Deviation 
(%)

	
	
	
	First
	Consequent
	

	4
	7
	2004
	48.69
	-
	

	
	
	2005
	
	47.87
	-1.68

	
	
	2006
	
	48.24
	-0.92

	
	
	2007
	
	47.81
	-1.81

	
	
	2008
	
	48.45
	-0.49

	
	
	2009
	
	48.20
	-1.01

	
	
	2010
	
	48.33
	-0.74


Results and Discussion

The results of µ ± SE and µ ± σN-1 for the ((%) are given in Figures 2 and 3 respectively. In general it can be seen that (a) all µ for the 29 ICs lie within the limit of ± 1.5%, which shows that the 29 ICs are stable in providing their ND,w, (b) the value of  µ ± SE  (Figure 2) is smaller than µ ± σN-1 (Figure 3) by a factor of square root of N, where N is the number of the deviations, as expected. For the results of IC=4 and RC=7 mentioned above, the µ ± SE and µ ± σN-1 are indicated by the arrows in the two figures.

We shall now examined µ ± SE only (Figure 2) for all the ICs as SE is in actual fact is the uncertainty in the mean. It can be seen that there are three ICs (IC no 7, 20 and 24) that give the values slightly higher than ± 1.5%. If we examine carefully the values of the SE for three ICs that lie outside acceptable limit of ± 1.5%, we found that their values are acceptable as they are less than 95% confidence interval of the limit, i.e.  79.6% or 1.27 standard deviation (SD) for IC no. 7 (1.91=1.27(1.5), 94.4% or 1.91 SD for IC no. 20 (2.86=1.91(1.5) and 75% or 1.15 SD for IC no. 24 (1.72=1.15(1.5). 
[image: image135.jpg]Percentage Deviation (%)

lower fimit = -1.5%

34567

89 10111213 14 1516 17 18 1920 2122 23 24 25 26 2728 29 30
ICNo




Figure 2. Deviation of µ ± SE
[image: image136.jpg]Percentage Deviation (%)

12345

678 9101112131415 16171819202122
ICNo

24252627

282930




Figure 3. Deviation of µ ± σN-1
Conclusion

As the ND,w calibration coefficient are stable with time, it is concluded that the chambers being studied are in their good performance for the purpose patients dose measurement. 
Acknowledgement
We gratefully acknowledge Universiti Kebangsaan Malaysia for a research grant, Project code UKM-ST-07-FRGS0154-2010.
References

1. IAEA. 1998. Building a better future: Contributions of nuclear science and technology. Vienna.

2. IAEA. 2004. Radiation, people and the environment. Vienna.

3. Samat, S.B., Evans, C.J., Kadni, T. Ooi, A.S. & Dollah, M.T. 2003. The -10% and 13% errors arising from mis-interpretation of the two IAEA protocols TRS-277 and TRS-398: Causes and prevention. SSDL Newsletter 48: 16-18.

4. IAEA. 2000. Absorbed Dose Determination in External Beam Radiotherapy: An International Code of Practice for Dosimetry Based on Standards of Absorbed Dose to Water, TRS 398, Vienna.

5. Samat, S.B., Evans, C.J., Kadni, T. & Dollah, M.T. 2009. Malaysian participation in the IAEA/WHO postal TLD and postal ionization chamber intercomparison programme: Analysis of results obtained during 1985-2008. Rad. Prot. Dos. 133(3): 186-191.

6. IAEA. 1994. Calibration of dosimeters used in radiotherapy: A manual sponsored by the IAEA and WHO. TRS 374. Vienna

7. IAEA. 2003. Results of comparison of therapy level ionization chamber calibration coefficient (Reference:MAL/02/01). IAEA/WHO Network of Secondary Standard Laboratories. 

142
147

