Dimple et al:  
DEVELOPMENT AND VALIDATION OF RP-HPLC METHOD FOR THE ESTIMATION OF CLOPIDOGREL BISULPHATE

Malaysian Journal of Analytical Sciences, Vol 17 No 3 (2013): 387 - 393

[image: image1.png]



DEVELOPMENT AND VALIDATION OF RP-HPLC METHOD FOR THE ESTIMATION OF CLOPIDOGREL BISULPHATE

(Pembangunan dan Penentusahkan Kaedah RP-HPLC Terhadap Klopidogrel Bisulfat)

Bhagat Dimple*, Mannur Vinodh, Mastiholimath Vinayak
Department of Quality Assurance, 

K.L.E .University’s College of Pharmacy, JNMC Campus,
 Nehrunagar, Belgaum 590 010, Karnataka, India
*Corresponding author: sukubhagat@gmail.com
Abstract

A novel, simple, precise, accurate and economical method has been developed and validated for the estimation of Clopidogrel bisulphate. Chromatographic analysis was performed on Hypersil BDS C18 (250 × 4.6 mm, 5 µm) column. Mobile phase used was potassium dihydrogen orthophosphate buffer: acetonitrile in a ratio of 32:68, pH was adjusted to 4.0 using orthophosphoric acid and 0.1 ml of triethyl amine was added for peak sharpness. Isocratic analysis was performed on SPD-20A double beam UV spectrophotometer at a detection range of 220 nm with a flow rate of 1.0 ml/min. The retention time for clopidogrel bisulphate was 3.847 min. The linear response was found to be in a concentration range of 50-150 µg/ml with a correlation coefficient of 0.999. The mean recovery was found to be 100.67 %. Intraday and Interday precision variations were found to be 1.88 and 0.863 respectively, which are within the limit of % RSD not more than 2. Hence the developed method is simple, fast, accurate and reproducible. It is suitable for routine quality control analysis of clopidogrel bisulphate. 

Keywords:  RP-HPLC, validation, clopidogrel bisulphate, linearity, accuracy, precision
Abstrak

Satu kaedah baru yang mudah, tepat, jitu dan ekonomikal telah dibangunkan dan disahkan terhadap analisa klopidogrel bisulfat. Analisis kromatografi telah dijalankan mengunakan turus Hypersil BDS C18 (250 × 4.6 mm, 5 μm ). Fasa bergerak yang digunakan ialah larutan penampan kalium dihidrogen ortofosfat : asetonitril dalam nisbah 32 : 68, pH telah diubah kepada nilai 4.0 menggunakan asid ortofosforik dan 0.1 ml larutan trietil amine ditambah untuk ketajaman puncak. Analisis isokratik telah dijalankan ke atas SPD- 20A spektrofotometer UV pada julat pengesanan 220 nm dengan kadar aliran 1.0 ml / min. Retensi masa bagi klopidogrel bisulfat adalah 3. 847 min. Sambutan linear didapati berada dalam julat kepekatan 50-150 μg ml dengan pekali korelasi adalah 0.999. Min perolehan semula adalah 100. 67 %. Variasi kejituan ujian Intra dan Inter-hari didapati masing – masing adalah 1.88 dan 0.863 yang berada dalam julat had % RSD tidak lebih daripada 2. Maka kaedah yang dibangunkan adalah mudah, cepat, tepat dan keboleh ulangan. Ia sesuai untuk analisis rutin kawalan mutu bagi  klopidogrel bisulfat.

Kata kunci: RP- HPLC, Penentusahkan, klopidogrel bisulfat., Kelinearan , Ketepatan, Kejituan
Introduction
Clopidogrel bisulphate is an antiplatelet agent or platelet aggregation inhibitor. The IUPAC name of clopidogrel bisulphate is Methyl (+)-(S)-α-(2-chlorophenyl)-6,7dihydrothieno [3,2-c]pyridine-5(4H)acetate sulfate (Figure 1).The drug is practically insoluble in water at neutral pH, but freely soluble at pH 1[1]. It is used to inhibit blood clots in coronary artery disease, peripheral vascular disease, and cerebrovascular disease [2]. 

[image: image2.jpg]« H,S0




Figure 1. Structure of clopidogrel bisulfate [1]

The quantitative analysis was made on Shimadzu prominence system using reverse phase chromatographic method to obtain less retention time. Hence clopidogrel bisulphate is insoluble in water; RP-HPLC method was used. Several analytical methods which are commonly used to identify the drug concentration and for quantification of drugs for the determination of clopidogrel bisulphate are Spectrofluorometry [3], LC-MS/MS [4], HPLC [5, 6], HPTLC [7, 8] and potentiometry method [9].
The aim of this study was to develop a RP-HPLC method, which could be employed for the routine analysis of the drug clopidogrel bisulphate.

Materials and Methods
Chemicals and Reagents 

Working standard of clopidogrel bisulphate was obtained from Sidmak India (pvt) Ltd Valsad, India having purity > 99.2%. HPLC grade Acetonitrile and Water was procured from Merck, Rankem Ltd., Mumbai, India. Phosphate buffer (pH 4.0) used as diluent and all chemical reagents used were HPLC grade.

Instrumentation 

Spectral scan was made on a Shimadzu UV 1800 pc double beam  spectrophotometer, with spectral bandwidth of 190-800 nm, wavelength accuracy of + 0.3 nm with automatic wavelength corrections using a pair of 10 mm quartz cells. All Spectral measurements were done using Spinchrom software. 
Chromatographic Conditions

Chromatographic analysis was performed on Thermo Hypersil BDS C18 (250 × 4.6 mm, 5 µm) column with a mobile phase used as potassium dihydrogen orthophosphate buffer adjusted pH 4.0 with orthophosphoric acid: acetonitrile in a ratio of 32:68 and 0.1 ml of triethyl amine was added for peak sharpness. The mobile phase was degassed and filtered through 0.45 μm membrane filter before pumping into HPLC system. Isocratic analysis was performed on SPD-20A double beam UV spectrophotometer at a detection range of 220 nm at a flow rate of 1.0 ml/min. The injection volume was 20 μl.

Preparation of Phosphate Buffer

Dissolve 10 g of Potassium dihydrogen orthophosphate and add 0.1ml of triethyl amine in 1000 ml distilled water. Adjust pH 4.0 with orthophosphoric Acid.
Preparation of Standard Solution

An accurately weighed quantity of 25mg clopidogrel bisulphate transferred in 25ml volumetric flask, dissolved with 10 ml of  acetonitrile and volume was made up to the mark with  acetonitrile sonicated for 10 min. From that solution pipette out 1 ml in 10 ml volumetric flask and volume was made up to the mark with mobile phase to make final Concentration 100 µg/ml.

Procedure
Equilibrate the column with mobile phase at prescribed condition until a stable baseline is achieved. Separately inject diluent, standard solution and test preparation into liquid chromatography. Measure the response for all peaks.
Results and Discussion
The developed method was validated for parameters like linearity, accuracy, precision, robustness, ruggedness, sensitivity and system suitability [10, 11]. 

Method development and optimization
Optimized chromatographic conditions for estimation of clopidogrel bisulphate are finalized using Column Thermo Hypersil BDS C18 (250×4.6 mm, 5 µm) consisting of mobile phase acetonitrile: phosphate buffer pH 4.0 in a ratio of 68:32. HPLC pump mode was made Isocratic with a flow rate of a 1.0 ml/min over the detection wavelength range of 220 nm. Standard spectrum of clopidogrel bisulphate is as shown in Figure 2.

Injection volume was made with Rheodyne injector 7725i of 20 µl. The run time was 10 min. The retention time for clopidogrel bisulphate was found to be 3.8 min.
[image: image3.png]amr

-




Figure 2. Chromatogram of standard preparation of Clopidogrel bisulphate

System Suitability

In these studies, 5 replicate injections of the drug at a concentration of 100 µg/ml were injected and evaluated for the parameters such as Retention time, theoretical plates and tailing factor.  The results obtained from the System Suitability studies were calculated and recorded in Table 1.
Table 1. System suitability parameters for Clopidogrel bisulphate

	Name
	Retention Time
	Area
	% Area
	Theoretical Plate Count
	Tailing Factor

	Clopidogrel bisulphate
	3.8
	1037.18
	100
	6073
	1.63


Linearity

The linear response was found in a concentration range of 50-150 µg/ml. The linearity plot is drawn in figure 3 and the correlation co-efficient (r2), y-intercept and slope of regression line were found to be 0.999, 9.521, -2.834 respectively and recorded in Table 2. 
[image: image4.png]1600
1400 =9.5215x- 2.8344
R?=0.9998

1200

1000
800 / 4 peakarea

600 ‘//’/ Linear (peak area)
400
200

0 50 100 150 200





Figure 3. Calibration curve of clopidogrel bisulphate at 220 nm.

Table 2. Linearity study for clopidogrel bisulphate

	Injection No. 
	Concentration (µg/ml)
	Area

	1
	50
	467.510

	2
	75
	715.577

	3
	100
	954.103

	4
	125
	1187.576

	5
	150
	1421.576


Accuracy

The accuracy was determined at three levels: 80, 100 and 120 %. From the peak area of the drug, percentage recovery (%) was calculated for 80, 100 and 120 %, which are 99.78, 101.54 and 100.70 % respectively. Percentage recovery was found to be in the limit of 98 – 102 % and recorded in Table 3. Hence, the method termed as accurate.
Precision
The precision expressed as standard deviation or Relative standard deviation. 
Intraday Precision
In this study, six injections of standard solution of  100 µg/ml were injected into the chromatographic system in different time intervals within a day. % RSD for retention time was found to be 0.768 and for area 1.88, which are within limit of % RSD not more than 2 and recorded in Table 4.

Table 3. Accuracy study for clopidogrel bisulphate

	Accuracy level
	Amount added 
	Area
	Amount recovered 
	% Recovery (Mean)
	SD
	% RSD

	80%
	80
	1708.768
	80.66
	99.78 %
	5.71
	0.336

	80%
	
	    1696.497
	79.39
	
	
	

	80%
	
	1696.798
	79.42
	
	
	

	100%
	100
	1913.944
	101.90
	101.54 %
	4.94
	0.258

	100%
	
	1903.472
	100.82
	
	
	

	100%
	
	1913.944
	101.90
	
	
	

	120%
	120
	2098.460
	121.00
	100.70 %
	11.02
	0.525

	120%
	
	2082.636
	119.36
	
	
	

	120%
	
	2109.498
	122.14
	
	
	


*average of three determinations

Table 4. Intraday precision of clopidogrel bisulphate
	Injection No.
	Retention Time (min)
	Area

	1
	3.810
	952.166

	2
	3.780
	950.675

	3
	3.780
	949.230

	4
	3.847
	975.080

	5
	3.853
	990.675

	6
	3.830
	950.788

	Average
	3.817
	959.135

	SD
	0.029
	18.05

	% RSD
	0.768
	1.88


Interday Precision

In this study, six injections of standard solution of 100 µg/ml were injected into the chromatographic system at different days. % RSD for retention time and for area was found to be 0.863 and 1.99 respectively, % RSD NMT 2 and recorded in Table 5.

Robustness

The robustness was obtained by altering the flow rate by 1.0 ml/min (altered flow rate 0.8 ml/min and 1.2 ml/min) pH 4.0 (altered pH 3.9 and 4.1). The results were obtained from robustness study were calculated. Percentage RSD for flow rate and for pH were found to be 1.49 and 1.03 respectively, % RSD NMT 2, and recorded in Table 6. Hence the method termed as robust.
Table 5. Interday Precision of Clopidogrel bisulphate

	Injection No.
	Retention Time
	Area

	1
	3.810
	952.166

	2
	3.780
	936.951

	3
	3.780
	949.230

	4
	3.853
	995.554

	5
	3.863
	1017.480

	6
	3.837
	970.751

	Average
	3.821
	969.355

	SD
	0.033
	28.96

	% RSD
	0.863
	1.99


Table 6. Robustness study of clopidogrel bisulphate

	Condition
	Modification
	Mean Area ± SD*
	% RSD
	Mean % RSD

	Flow Rate (ml/min)
	0.8 ml/min
	1041.542 ± 17.83
	1.71
	1.49

	
	1.2 ml/min
	902.667 ± 11.43
	1.27
	

	pH
	3.9
	952.471 ± 7.63
	0.80
	1.03

	
	4.1
	968.01 ± 12.18
	1.26
	


*average of three determinations

Ruggedness
The ruggedness of the method was determined by carrying out the experiment by different operators. Percentage RSD were calculated and recorded in Table 7. Result obtained by different analyst and different day were found within acceptable limit, which shows that test method is rugged.
Table 7. Ruggedness study of Clopidogrel bisulphate

	Injection no.
	           Analyst Ι
	         Analyst ΙΙ

	
	          Area
	
	             Area
	

	1
	          479.040
	
	            480.981
	

	2
	          486.911
	
	            493.290
	

	Average
	          482.976
	            487.136

	SD
	              3.940
	                6.150

	% RSD
	              0.810
	                1.260

	
	


Limit of Detection and Quantification (LOD and LOQ)

The Limit of Detection and Quantification, LOD and LOQ data was calculated and recorded in Table 8. LOD and LOQ were found to be 1.8728, 5.6752 ppm respectively.
Table 8. Limit of Detection and Quantification in this study
	Parameters
	Clopidogrel bisulphate

	LOD (µg/ml)
	1.8728

	LOQ (µg/ml)
	5.6752


Conclusion
The developed analytical method is validated as per ICH guidelines. A rapid and simple RP-HPLC method for determination of clopidogrel bisulphate has been developed and validated. The method involved a mobile phase consisting of acetonitrile: phosphate buffer (pH 4.0) 68:32 used at 220 nm. The retention time was 3.837 min at a flow-rate of 1.0 ml/min and the injection volume was 20 μl. The total run time for an assay was approximately 10 min. This proposed method was validated for parameters including linearity, accuracy, sensitivity, precision, ruggedness and robustness. The present study is concerned for the development and validation of simple, specific, sensitive, accurate and precise RP-HPLC method for the determination of clopidogrel bisulphate for routine analysis work.
Acknowledgement
The authors wish to thank Dr. Anjan Paul from Sidmak India (pvt) Valsad, Gujarat for providing drug sample and Mr. Ketan Patel from Molecule Laboratory, Ahmadabad, Gujarat for providing the facilities to carry out this research work.

References
1. Clopidogel Drug Description on www.rxlist.com/plavix-drug.htm.

2. Clopidogrel Information on en.wikipedia.org/ wiki/ Clopidogrel.
3. Saddet, D., & Erkem, A. (2009). Rapid and accurate determination of clopidogrel in tablets by using spectrophotometric and chromatographic techniques. Commun Fac Sci Univ Ank Series B, 55 (1) : 1-16.

4. Venkanna, B., Shreedhara, C., & Ajitha, M. (2011). Rapid and rugged bio analytical method development and validation clopidogrel in human plasma using liquid chromatography/ tandem mass spectroscopy, American Journal Pharmtech Research, 1 (2) : 66 - 80.

5. Sahu, S., & Sarangi, S. Development and validation method for the estimation of clopidogrel in bulk and pharmaceutical dosage form, International Journal Research of Pharmaceutical Science, 3 (2) : 224-27.

6. Josh, K., & Jayasehar, P. (2013). Stability indicating HPLC determination of clopidogrel bisulfate in pharmaceutical dosage forms, International Journal of Comprehensive Pharmacy, 1 (06): 1 - 5. 

7. Himani, A., Kaul, N., Paradkar, A.R., Mahadik, K.R. (2003). Stability indicating HPTLC determination of clopidogrel bisulphate as bulk drug and in pharmaceutical dosage form, Talanta 61: 581-89.

8. Sinha, P.K., & Mrinalini, C. (2009). A Validated Stability Indicating HPTLC Method for Determination of Aspirin and Clopidogrel Bisulphate in Combined Dosage Form, Eurasian J Anal Chem, 4 (2) : 152-60.

9. SABER, L.A., Elmosallamy, M.A. (2008). Journal of Food and Drug Analysis, Liquid Chromatographic and Potentiometric Methods for Determinations of Clopidogrel, 16 (2): 11-18.

10. Validation of Analytical Procedures: Methodology, ICH Harmonised Tripartite Guidelines, (1996), 1-11.

11. Validation of Analytical Procedures: Methodology, ICH Harmonised Tripartite Guidelines, (1994), 1-7.

392
393

